Supplementary Data: DOI: 10.1515/cclm-2015-0578

Supplementary Table 1: LC-method. The following eluents were used: 10 mmol/L ammonium acetate in water + 0.1% (v/v) formic acid (eluent A), 10 mmol/L ammonium acetate in methanol/acetonitrile 50/50 v/v + 0.1% (v/v) formic acid (eluent B), acetone/acetonitrile/2-propanol 1/1/1 v/v/v (eluent C). EP, eluting pump (analytical chromatography); LP, loading pump (online extraction).

| Time (min) | Loading pump flow rate (ml/min) | %A | %B | %C | Eluting pump flow rate (ml/min) | %A | %B | %C |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 0 | 2 | 99 | 1 | - | 0.3 | 99 | 1 | - |
| 0.50 | 0.2 | 99 | 1 | - | 0.3 | 99 | 1 | - |
| 1.33 | 2 | - | - | 100 | 0.3 | 50 | 50 | - |
| 3.67 | 1 | 60 | 40 | - | 0.3 | 50 | 50 | - |
| 5.33 | 2 | 60 | 40 | - | 0.4 | - | - | 100 |
| 6.17 | 2 | 60 | 40 | - | 0.3 | - | 100 | - |
| 6.50 | 2 | 99 | 1 | - | 0.3 | 99 | 1 | - |
| 8.50 | 2 | 99 | 1 | - | 0.3 | 99 | 1 | - |