
Supplemental Material: https://doi.org/10.1515/cclm-2016-1183

Supplemental Table  1 :  FT4 (pmol/L), TSH (mUI/L) and PTH (ng/L) values obtained with alternative methods (left), and after neutralisation (protocol N5), with the Roche methods (right).  Values in bold are outside of the method related reference range. 3 subjects were sampled twice (T1 same subject as T2, T3 same subject as  T4, T7 same subject as T8).
	Samples
	biotin
	FT4
	TSH
	PTH
	
	FT4
	TSH
	PTH

	
	µg/L
	Liaison
	Access
	Liaison
	
	Roche (after biotin neutralization)

	P1
	507
	14.4
	0.699
	24.5
	
	17.6
	0.827
	22.9

	P2
	690
	14.3
	1.06
	24.8
	
	17.5
	1.23
	25.0

	P3
	487
	12.9
	1.28
	30.4
	
	14.6
	1.49
	30.5

	P4
	694
	15.2
	2.24
	86.9
	
	16.2
	1.75
	62.1

	P5
	180
	13.6
	1.0
	26.2
	
	21.4
	1.05
	28.7

	P6
	363
	15.6
	1.64
	22.2
	
	18.6
	1.87
	20.3

	P7
	169
	11.0
	3.42
	29.3
	
	16.7
	3.87
	26.3

	P8
	233
	13.8
	 
	31.1
	
	17.7
	1.46
	34.7

	P9
	176
	 
	 
	52.1
	
	15.3
	2.4
	39.9

	T1
	827
	13.1
	 
	31.2
	
	15.0
	1.4
	33.5

	T2
	758
	 
	 
	31.3
	
	15.1
	1.31
	31.5

	T3
	1000
	12.1
	1.71
	30.1
	
	18.8
	0.966
	26.3

	T4
	1110
	 
	 
	 
	
	16.6
	2.0
	23.1

	T5
	1160
	11.2
	 
	25.5
	
	13.3
	2.39
	26.0

	T6
	669
	13.2
	 
	20.4
	
	16.4
	2.29
	19.0

	T7
	1010
	13.4
	1.66
	31.6
	
	15.8
	1.3
	26.1

	T8
	959
	13.4
	 
	25.0
	
	15.6
	1.13
	29.9

	T9
	407
	13.2
	0.945
	25.4
	
	16.3
	1.21
	22.5

	T10
	914
	 13.8
	1.55
	 
	
	16.7
	1.32
	18.2

	T11
	814
	11.8
	 
	24.6
	
	14.1
	0.733
	18.8

	V1
	43.9
	16.1
	1.32
	38.6
	
	17.6
	1.54
	41.9

	V2
	56.8
	14.9
	1.20
	58.6
	
	18.5
	1.07
	54.5

	V3
	31.7
	17.9
	2.50
	72.5
	
	16.9
	2.87
	56.6


Supplemental Table  2 :  Hormonal results in 39 biotin-unsupplemented patients, before and after neutralization (protocol N5, in parenthesis) (thyroid function test n= 12; cortisol n=6; FSH, LH and Prolactin n=10; PTH and 25OHD n= 11). CV : maximum CV, from daily internal quality control database.

	
	PTH, 

ng/L
	25OHD, µg/L
	FSH, UI/L
	LH, UI/L
	PROL, µg/L
	CORTISOL,

µg/dl
	FT3, pmol/L
	FT4, pmol/L
	TSH, mUI/L

	
	71.5 (70.5)
	24.7 (25.3)
	3.3 (3.5)
	0.71 (0.77)
	20.4 ( 20.7)
	0.46 (0.43)
	2.7 (2.6)
	12.5 (12.6)
	1.11 ( 1.09)

	
	30.8 (29.6)
	40.7 ( 43.5)
	4.0 (4.0)
	3.6 ( 3.7)
	8.0 (7.9)
	13.1  (12.5)
	4.3 (4.3)
	13.8 (13.6)
	0.021 (0.021)

	
	93.4 (83.2)
	25.1 (24.7)
	1.2 (1.2)
	<0.1 (<0.1)
	13.0 (12.9)
	5.98 ( 5.83)
	2.8 (2.7)
	38.0 (39.5)
	0.018 (0.017)

	
	69.7 (63.3)
	13.6 ( 13.0)
	6.3 (6.7)
	1.4 (1.5)
	15.9 (16.5)
	30.4 (28.7)
	1.8 (1.7)
	15.4 (15.4)
	16.3 (16.4)

	
	788 (696)
	16.6 (16.6)
	2.2 (2.2)
	0.14 (0.14)
	11.0 (10.8)
	0.91 (0.90)
	4.7 (4.9)
	22.7 (23.3)
	0.101 (0.099)

	
	162 (150)
	44.1 (45.6)
	6.3 (6.8)
	3.1 (3.2)
	39.4 (40.2)
	1.10 (1.04)
	4.6 (4.8)
	10.4 (10.7)
	47.2 (47.0)

	
	81.2 (68.7)
	11.0 (11.0)
	4.8 (4.9)
	3.4 (3.7)
	55.9 (58.6)
	
	2.7 (2.6)
	9.7 (9.9)
	0.77 (0.75)

	
	277 ( 242)
	13.9 (12.5)
	0.1 (0.1)
	<0.1 (<0.1)
	1.6 (1.6)
	
	3.7 (3.6)
	20.5 (20.4)
	0.048 (0.051)

	
	7.1 (7.5)
	22.0 (18.9)
	4.7 (4.6)
	2.9 (2.9)
	19.8 (20.2)
	
	4.3 (3.9)
	32.8 (32.8)
	0.005 (0.005)

	
	119 (100)
	46.4 (45.0)
	98.3 (101)
	38.6 (39.6)
	35.8 (35.3)
	
	5.0 (5.0)
	32.1 (32.7)
	0.107 (0.109)

	
	32.6 (29.4)
	45.2 (44.4)
	
	
	
	
	4.1 (4.0)
	19.1 (19.3)
	56.0 (54.1)

	
	
	
	
	
	
	
	6.7 (6.6)
	20.4 (20.7)
	0.055 (0.057)

	CV, %
	5.1
	8.8
	4.5
	4.8
	4.4
	3.5
	4.6
	3.6
	4.2

	Mean difference

after N5, %
	-8.4
	-2.1
	2.0
	3.9
	0.8
	-2.3
	-1.8
	1.2
	0.1

	Maximum difference 

after N5, %
	-15.7
	6.8
	7.8
	8.9
	4.8
	-5.6
	3.5
	3.9
	6.2




Page 1/1

