Supplemental files to the paper “Performance of Afinion HbA1c measurements in General Practice as judged by External Quality Assurance data”, by Stavelin A, et. al.

Figure 1. Youden plots for each HbA1c survey showing Afinion participant result 1 (control sample 1) vs Afinion participant result 2 (control sample 2) in relation to target value (red lines) and the participant performance specifications of ± 6% (green lines).

Afinion HbA1c reagent lot numbers

Table 1-7 show the number of participants, the target value and median result for each kit reagent lot number for control sample 1 and 2 in the surveys 1-2017, 2-2017, 3-2017, 4-2017, 1-2018, 2-2018 and 3-2018.

Each lot number represent a unique lot code, meaning that if the lot code is the same in two surveys, the lot number is also the same (e.g. lot code 9 represent lot number 10185511 in survey 1-2017, 2-2017 and 3-2017).

Table 1. Afinion reagent lot evaluation in Noklus EQA HbA1c survey 1-2017.

Table 2. Afinion reagent lot evaluation in Noklus EQA HbA1c survey 2-2017.

Table 3. Afinion reagent lot evaluation in Noklus EQA HbA1c survey 3-2017.

| Target level |  |  | Control sample 1 |  | Control sample 2 |  |
| --- | --- | --- | --- | --- | --- | --- |
|  |  |  | 5.25 |  | 6.44 |  |
| Lot code | Kit reagent lot number | Exp. date | Median | N | Median | N |
| 4 | 10182328 | 31.03.2018 | 5.25 | 7 | 6.45 | 7 |
| 9 | 10185511 | 31.10.2018 | 5.15 | 10 | 6.33 | 10 |
| 10 | 10185926 | 30.11.2018 | 5.15 | 14 | 6.35 | 15 |
| 11 | 10186147 | 30.11.2018 | 5.25 | 21 | 6.45 | 20 |
| 12 | 10186921 | 31.12.2018 | 5.25 | 90 | 6.45 | 89 |
| 13 | 10187455 | 28.02.2019 | 5.25 | 13 | 6.35 | 13 |
| 14 | 10188397 | 31.03.2019 | 5.15 | 117 | 6.35 | 117 |
| 15 | 10188585 | 31.03.2019 | 5.15 | 90 | 6.35 | 90 |
| 16 | 10189349 | 31.05.2019 | 5.20 | 296 | 6.45 | 292 |
| 17 | 10189769 | 30.06.2019 | 5.25 | 16 | 6.40 | 16 |
| 18 | 10190252 | 30.06.2019 | 5.25 | 13 | 6.43 | 12 |

Table 4. Afinion reagent lot evaluation in Noklus EQA HbA1c survey 4-2017.

| Target level |  |  | Control sample 1 |  | Control sample 2 |  |
| --- | --- | --- | --- | --- | --- | --- |
|  |  |  | 6.12 |  | 6.98 |  |
| Lot code | Kit reagent lot number | Exp. date | Median | N | Median | N |
| 12 | 10186921 | 31.12.2018 | 6.10 | 33 | 7.05 | 33 |
| 13 | 10187455 | 28.02.2019 | 6.05 | 9 | 6.90 | 9 |
| 14 | 10188397 | 31.03.2019 | 5.95 | 17 | 6.85 | 17 |
| 15 | 10188585 | 31.03.2019 | 6.00 | 13 | 6.85 | 13 |
| 16 | 10189349 | 31.05.2019 | 6.05 | 50 | 6.90 | 49 |
| 17 | 10189769 | 30.06.2019 | 6.00 | 7 | 6.90 | 7 |
| 18 | 10190252 | 30.06.2019 | 6.05 | 273 | 6.90 | 270 |
| 19 | 10190590 | 31.08.2019 | 6.05 | 289 | 6.95 | 290 |
| 20 | 10190990 | 30.09.2019 | 6.05 | 8 | 6.88 | 8 |

Table 5. Afinion reagent lot evaluation in Noklus EQA HbA1c survey 1-2018.

| Target level |  |  | Control sample 1 |  | Control sample 2 |  |
| --- | --- | --- | --- | --- | --- | --- |
|  |  |  | 5.01 |  | 6.24 |  |
| Lot code | Kit reagent lot number | Exp. date | Median | N | Median | N |
| 11 | 10186147 | 30.11.2018 | 5.05 | 5 | 6.25 | 5 |
| 12 | 10186921 | 31.12.2018 | 5.05 | 13 | 6.20 | 12 |
| 13 | 10187455 | 28.02.2019 | 5.00 | 6 | 6.13 | 6 |
| 14 | 10188397 | 31.03.2019 | 4.90 | 9 | 6.05 | 9 |
| 15 | 10188585 | 31.03.2019 | 4.95 | 8 | 6.08 | 8 |
| 16 | 10189349 | 31.05.2019 | 4.90 | 22 | 6.15 | 22 |
| 18 | 10190252 | 30.06.2019 | 5.00 | 74 | 6.10 | 74 |
| 19 | 10190590 | 31.08.2019 | 4.95 | 100 | 6.10 | 99 |
| 20 | 10190990 | 30.09.2019 | 4.90 | 97 | 6.05 | 97 |
| 21 | 10193607 | 31.10.2019 | 4.95 | 347 | 6.10 | 345 |

Table 6. Afinion reagent lot evaluation in Noklus EQA HbA1c survey 2-2018.

| Target value -> |  |  | Control sample 1 |  | Control sample 2 |  |
| --- | --- | --- | --- | --- | --- | --- |
|  |  |  | 5.48 |  | 7.72 |  |
| Lot code | Lot number | Exp. Date | Median | n | Median | n |
| 16 | 10189349 | 05.2019 | 5.40 | 8 | 7.68 | 8 |
| 18 | 10190252 | 06.2019 | 5.40 | 15 | 7.60 | 15 |
| 19 | 10190590 | 08.2019 | 5.35 | 17 | 7.65 | 17 |
| 20 | 10190990 | 09.2019 | 5.25 | 18 | 7.50 | 18 |
| 21 | 10193607 | 10.2019 | 5.35 | 243 | 7.60 | 240 |
| 22 | 10194033 | 11.2019 | 5.40 | 176 | 7.65 | 176 |
| 23 | 10194835 | 01.2020 | 5.45 | 174 | 7.65 | 174 |
| 24 | 10195320 | 02.2020 | 5.35 | 18 | 7.65 | 18 |

Table 7. Afinion reagent lot evaluation in Noklus EQA HbA1c survey 3-2018.

| Target value -> |  |  | Control sample 1 |  | Control sample 2 |  |
| --- | --- | --- | --- | --- | --- | --- |
|  |  |  | 5.00 |  | 6.76 |  |
| Lot code | Lot number | Exp. Date | Median | n | Median | n |
| 18 | 10190252 | 06.2019 | 4.88 | 8 | 6.70 | 8 |
| 19 | 10190590 | 08.2019 | 5.00 | 7 | 6.80 | 7 |
| 21 | 10193607 | 10.2019 | 4.95 | 42 | 6.70 | 41 |
| 22 | 10194033 | 11.2019 | 5.00 | 21 | 6.70 | 22 |
| 23 | 10194835 | 01.2020 | 5.05 | 35 | 6.75 | 35 |
| 24 | 10195320 | 02.2020 | 4.95 | 110 | 6.75 | 112 |
| 25 | 10195436 | 02.2020 | 4.95 | 67 | 6.75 | 69 |
| 26 | 10195679 | 03.2020 | 4.90 | 239 | 6.80 | 235 |
| 27 | 10196709 | 05.2020 | 5.00 | 148 | 6.75 | 148 |
| 28 | 10197038 | 05.2020 | 5.05 | 7 | 6.75 | 7 |