Supplemental Tables and Figures

Table 1. Precursor and product transitions and mass spectrometer settings.

|  | Precursor ion, m/z | Product ion, m/z |  | Collision energy, eV | Cone voltage, V |
| --- | --- | --- | --- | --- | --- |
| Analytes |  |  |  |  |  |
| 6-Sulfatoxymelatonin (quantifier) 6-Sulfatoxymelatonin (qualifier) | 329 329 | 190 158 |  | 23 35 | 22 22ulfatoxymelatoninopey |
| Internal standards |  |  |  |  |  |
| 6-Sulfatoxymelatonin-d4 (quantifier) 6-Sulfatoxymelatonin-d4 (qualifier) | 333 333 | 194 161 |  | 23 35 | 22 22 |

m/z, mass over charge ratio; eV, electron volt; V, volt.

Table 2. Intra- and interassay imprecision and recovery results.

|  | Intraassaya |  |  | Interassayb |  |  | Recovery |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  | Mean, nmol/L | SD, nmol/L | CV, % | Mean, nmol/L | SD, nmol/L | CV, % | Mean (range) % |
| Low | 6.39 | 0.16 | 2.5 | 6.43 | 0.35 | 5.4 | 97.4 (90–104) |
| Medium | 30.1 | 0.60 | 2.0 | 30.7 | 0.90 | 2.9 | 101 (90–111) |
| High | 111 | 2.4 | 2.1 | 114 | 3.8 | 3.4 | 101 (91–115) |
| a n=20; b n=12. |  |  |  |  |  |  |  |

CV, coefficient of variation; SD, standard deviation.

Table 3. Descriptive data of study population in the biological variation study.

| Individual | Sex | Age, years |
| --- | --- | --- |
| 1 | ♂ | 24 |
| 2 | ♀ | 58 |
| 3 | ♀ | 29 |
| 4 | ♀ | 33 |
| 5 | ♂ | 39 |
| 6 | ♂ | 19 |
| 7 | ♀ | 21 |
| 8 | ♀ | 21 |
| 9 | ♀ | 48 |
| 10 | ♀ | 22 |

Figure 1. Scatter plot to compare two methods for measuring 6-sulfatoxymelatonin in 40 urine samples.

X-axis: 6-sulfatoxymelatonin concentrations measured by ELISA (in nmol/L). Y-axis: 6-sulfatoxymelatonin concentrations measured by LC-MS/MS (in nmol/L). Dashed line represents the line of identity. Passing-Bablok regression analysis is shown by the solid line. Blue shaded area represents the 95% confidence interval. Dashed line is the line of identity. Abbreviations: 6-SM, 6-sulfatoxymelatonin; LC-MS/MS, liquid chromatography tandem mass spectrometry.

Figure 2. Scatter plot of urinary 6-sulfatoxymelatonin excretion (nmol/24 h) in healthy individuals between 20 and 80 years old.