Table S1: Repeatability and within-laboratory precision of SARS-CoV-2 IgM

|  |  | Repeatability |  | Within-Laboratory |  |
| --- | --- | --- | --- | --- | --- |
| Sample | Mean, kAU/L | SD, kAU/L | CV, % | SD, kAU/L | CV, % |
| Serum 1 | 7.29 | 0.31 | 4.32% | 0.32 | 4.45% |
| Serum 2 | 13.5 | 0.41 | 3.06% | 0.44 | 3.26% |
| Serum 3 | 98.2 | 2.75 | 2.80% | 2.97 | 3.02% |
| Serum 4 | 295 | 8.93 | 3.03% | 9.07 | 3.08% |

Table S2: Repeatability and within-laboratory precision of SARS-CoV-2 IgG

|  |  | Repeatability |  | Within-Laboratory |  |
| --- | --- | --- | --- | --- | --- |
| Sample | Mean, kAU/L | SD, kAU/L | CV, % | SD, kAU/L | CV, % |
| Serum 1 | 7.44 | 0.32 | 4.30% | 0.32 | 4.30% |
| Serum 2 | 13.6 | 0.42 | 3.11% | 0.42 | 3.12% |
| Serum 3 | 77.8 | 2.54 | 3.26% | 3.99 | 5.13% |