Supplemental material 3

|  |  |  |  |  |  | Courier |  |  |  | PTS | Difference (PTS-courier) | Difference (PTS/courier, %) |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Analyte | Level (mean) | unit | N | LoD | Upper limit of quantitation | Minimum | Median | Maximum | Mean | Mean | Mean (95%CI) | Lower Limit of Agreement (95%CI) | Upper Limit of Agreement (95%CI) |
| Albumin | <20 | mg/L | 54 | 3 | 4400 | 3.00 | 3.5 | 18.2 | 6.16 | 6.01 | -0.15 (-0.344 to 0.0437) | -19.7 (-23.5 to -15.7) | 20.1 (14.4 to 26.1) |
| Albumin | >/=20 | mg/L | 46 | 3 | 4400 | 19.9 | 181 | 2390 | 470 | 470 | 3.23 (-2.85 to 9.32) | -3.92 (-4.96 to -2.86) | 4.44 (3.31 to 5.59) |
| Albumin-to-Creatinine Ratio | <10 | mg/mmol | 64 | N/A | N/A | 0.113 | 1.02 | 8.60 | 1.71 | 1.73 | 0.0141 (-0.0278 to 0.0561) | -18.3 (-21.8 to -14.6) | 21.6 (16.4 to 27.0) |
| Albumin-to-Creatinine Ratio | >/=10 | mg/mmol | 36 | N/A | N/A | 11.9 | 53.7 | 437 | 97.3 | 97.8 | 0.506 (-0.806 to 1.88) | -4.99 (-6.48 to -3.47) | 5.61 (3.96 to 7.30) |
| Creatinine | <5 | mmol/L | 41 | 0.1 | 135 | 0.87 | 3.07 | 5 | 3.16 | 3.15 | -0.0132 (-0.0387 to 0.0123) | -6.51 (-8.11 to -4.89) | 5.76 (3.96 to 7.6) |
| Creatinine | >/=5 | mmol/L | 59 | 0.1 | 135 | 5.33 | 8.34 | 34.1 | 10.7 | 10.7 | -0.0336 (-0.103 to 0.0359) | -4.82 (-5.85 to -3.77) | 4.64 (3.5 to 5.79) |
| Potassium | All | mmol/L | 99 | 3 | 100* | 6.34 | 36.6 | 172 | 40.4 | 40,5 | 0.114 (-0.276 to 0.505) | -8.35 (-9.71 to -6.97) | 8.62 (7.00 to 10.3) |
| Protein | <0.15 | g/L | 63 | 0 | 6000 | 0.04 | 0.06 | 0.14 | 0.07 | 0.07 | 0.000048 (-0.00103 to 0.00112) | -11.5 (-13.9 to -9.11) | 12.7 (9.67 to 15.7) |
| Protein | >/=0.15 | g/L | 37 | 0 | 6000 | 0.15 | 0.35 | 2.94 | 0.76 | 0.76 | -0.00297 (-0.00866 to 0.00271) | -5.14 (-6.51 to -3.74) | 4.78 (3.25 to 6.32) |
| Sodium | <100 | mmol/L | 66 | 20 | 350 | 20.0 | 62.9 | 99.7 | 61.2 | 62,4 | 1.17 (-0.307 to 2.64) | -17.1 (-20.8 to -13.2) | 26.4 (20.7 to 32.3) |
| Sodium | >/=100 | mmol/L | 34 | 20 | 350 | 100 | 140 | 260 | 150 | 150 | 0.271 (-0.903 to 1.44) | -4.88 (-6.42 to -3.32) | 5.75 (4.03 to 7.49) |
| Urea Nitrogen | All | mmol/L | 100 | 1 | 3600 | 27.9 | 160 | 516 | 178 | 178 | 0.0470 (-0.808 to 0.902) | -5.42 (-6.34 to -4.49) | 5.75 (4.72 to 6.79) |

Table S3 - A: Soluble components. * Manual dilution of samples for extended upper limit of quantification of 1000 mmol/L.

|  |  |  |  |  |  | Courier |  |  |  | PTS | Difference (PTS-courier) | %Difference (PTS/courier) |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Analyte | Level (mean) | unit | N | LoD | Upper limit of quantitation | Minimum | Median | Maximum | Mean | Mean | Mean (95%CI) | Lower Limit of Agreement (95%CI) | Upper Limit of Agreement (95%CI) |
| Bacteria | 5-30 | x10^6/L | 25 | 5 | 10000 | 25 | 3.4 | 32 | 15 | 16 | 1.2 (-1.2 to 3.7) | -57 (-69 to -41) | 150 (82 to 250) |
| Bacteria | 30-1000 | x10^6/L | 31 | 5 | 10000 | 41 | 150 | 1000 | 250 | 260 | 8.8 (-7.5 to 25) | -19 (-24 to -13) | 26 (17 to 35) |
| Bacteria | 1000-10000 | x10^6/L | 14 | 5 | 10000 | 1100 | 2300 | 6600 | 3200 | 3200 | 78 (-37 to 190) | -11 (-17 to -5.6) | 14 (6.7 to 21) |
| Crystals | All | x10^6/L | 1 | 10 | 200 | 54 | 54 | 54 | 54 | 51 | -3.4 |  |  |
| Hyaline Casts | All | x10^6/L | 6 | 1 | 30 | 0.82 | 1.9 | 6.3 | 2.6 | 3.1 | 0.54 (-0.75 to 1.8) | -36 (-64 to -12) | 120 (23 to 280) |
| Non-lysed Red Blood Cells | 1-10 | x10^6/L | 64 | 1 | 10000 | 0 | 2.6 | 8.5 | 3.2 | 4.1 | 0.91 (0.42 to 1.4) | -62 (-71 to -50) | 370 (260 to 510) |
| Non-lysed Red Blood Cells | 10-50 | x10^6/L | 14 | 1 | 10000 | 12 | 17 | 48 | 20 | 21 | 0.39 (-1 to 1.8) | -22 (-32 to -11) | 34 (16 to 53) |
| Non-lysed Red Blood Cells | >/=50 | x10^6/L | 8 | 1 | 10000 | 60 | 160 | 2000 | 470 | 480 | 10 (-6.5 to 27) | -6.6 (-11 to -1.7) | 7.2 (1.9 to 13) |
| Pathological casts | All | x10^6/L | 1 | 1 | 30 | 2.51 | 2.51 | 2.51 | 2.51 | 4.74 | 2.23 |  |  |
| Red Blood Cells | 1-10 |  | 71 | 1 | 10000 | 0 | 3 | 9.3 | 3.5 | 4.6 | 1.1 (0.58 to 1.6) | -60 (-69 to -48) | 380 (270 to 530) |
| Red Blood Cells | 10-50 | x10^6/L | 14 | 1 | 10000 | 14 | 18 | 37 | 20 | 21 | 1.2 (-0.68 to 3.1) | -23 (-35 to -9.7) | 47 (24 to 73) |
| Red Blood Cells | >50 | x10^6/L | 9 | 1 | 10000 | 52 | 120 | 2100 | 440 | 450 | 11 (-6.3 to 28) | -6.6 (-11 to -2) | 7.6 (2.6 to 13) |
| Renal Tubular Epithelial Cells | <2.5 | x10^6/L | 24 | 1 | 200 | 0.70 | 1.3 | 2.5 | 1.3 | 1.8 | 0.44 (0.14 to 0.73) | -44 (-59 to -23) | 200 (120 to 310) |
| Renal Tubular Epithelial Cells | >/=2.5 | x10^6/L | 27 | 1 | 200 | 2.8 | 5.5 | 19 | 6.2 | 5.7 | -0.46 (-1.1 to 0.15) | -43 (-52 to -32) | 54 (29 to 83) |
| Spermatozoa | All | x10^6/L | 0 | 50 | 200 |  |  |  |  |  |  |  |  |
| Squamous Epithelial Cells | <10 | x10^6/L | 25 | 1 | 200 | 1 | 2.9 | 9.8 | 3.9 | 3.9 | -0.048 (-0.5 to 0.41) | -45 (-55 to -32) | 71 (39 to 110) |
| Squamous Epithelial Cells | >/= 10 | x10^6/L | 32 | 1 | 200 | 10 | 23 | 150 | 35 | 33 | -1.1 (-3.4 to 1.1) | -23 (-29 to -17) | 26 (16 to 36) |
| Transitional Epithelial Cells | All | x10^6/L | 11 | 1 | 200 | 0.4 | 1.3 | 37 | 5.3 | 5.4 | 0.13 (-1.3 to 1.5) | -74 (-89 to -40) | 330 (88 to 900) |
| White Blood Cells | 1-15 | x10^6/L | 50 | 1 | 10000 | 1 | 3.5 | 13 | 4.6 | 5.2 | 0.54 (-0.067 to 1.2) | -51 (-60 to -40) | 140 (100 to 200) |
| White Blood Cells | 15-50 | x10^6/L | 16 | 1 | 10000 | 16 | 25 | 41 | 26 | 26 | -0.16 (-2 to 1.7) | -27 (-37 to -17) | 30 (14 to 50) |
| White Blood Cells | 50-1000 | x10^6/L | 15 | 1 | 10000 | 55 | 120 | 380 | 170 | 170 | 4.8 (-8.3 to 18) | -16 (-23 to -9) | 19 (9 to 29) |
| White Blood Cells | >/=1000 | x10^6/L | 4 | 1 | 10000 | 1000 | 5600 | 16000 | 7200 | 7000 | -150 (-420 to 130) | -4.6 (-8.6 to -0.34) | 1.5 (-2.8 to 6) |
| Yeast-like Cells | <4 | x10^6/L | 21 | 1 | 200 | 0.0 | 1.3 | 3.4 | 1.4 | 2.0 | 0.55 (0.13 to 0.96) | -48 (-65 to -22) | 270 (150 to 450) |
| Yeast-like Cells | >/=4 | x10^6/L | 3 | 1 | 200 | 4.1 | 32 | 330 | 120 | 130 | 8.9 (-37 to 55) | -85 (-100 to 3700) | 2200 (-91 to 570000) |

Table S3-B: particles.