| Analyte | Units | Analytical Imprecision |  |  | LLOQ | Reference |
| --- | --- | --- | --- | --- | --- | --- |
|  |  | Control Level 1 Concentration SD CV (%) | Control Level 2 Concentration SD CV (%) | Control Level 3 Concentration SD CV (%) |  |  |
| Alanine Aminotransferase - Without Piridoxal Phosphate (ALT) | U/L | 33 1.6 4.8 | 90 2.6 2.9 | --- | 7 | 35 |
| Alkaline Phosphatase, 4-Nitrophenyl Phosphate (ALP) | U/L | 119 7.1 6.0 | 434 9.9 2.3 | --- | 9 | 35 |
| Cholesterol | mmol/L | 6.45 0.186 2.9 | 2.89 0.084 3.1 | --- | 5 | 35 |
| Creatinine (Jaffe) | μmol/L | 127 3.7 3.0 | 579 12.4 2.1 | --- | 8.8 | 35 |
| Creatinine (Enzymatic) | μmol/L | 43 1.3 4.0 | 490 6.2 1.3 | --- | 8.8 | 35 |
| Direct HDL-Cholesterol (UHDL) | mmol/L | 1.51 0.039 2.6 | 0.69 0.21 3.0 | --- | 0.13 | 35 |
| Estradiol | pmol/L | 29.4 3.22 10.9 | 230.8 7.33 3.2 | 498.2 28.1 5.6 | -- | 34 |
| Ferritin | µg/L | 30.3 1.87 6.2 | 210.6 14.86 7.1 | 476.4 32.55 6.8 | -- | 33 |
| FSH | IU/L | 7.20 0.49 6.8 | 17.04 1.40 8.2 | 39.41 2.77 7.0 | -- | 34 |
| Iron | μmol/L | 31.0 1.21 3.9 | 5.1 0.45 8.9 | --- | 1.3 | 35 |
| LH | IU/L | 3.43 0.34 9.9 | 14.89 146 9.8 | 52.88 3.62 6.9 | 0.09 | 34 |
| Testosterone | nmol/L | 8.93 0.692 7.9 | 74.93 5.13 6.9 | 219.88 14.32 6.5 | 0.15 | 34 |
| Triglycerides | mmol/L | 1.37 0.057 4.1 | 5.02 0.064 1.3 | --- | 0.071 | 35 |
| Vitamin B12 | pmol/L | 270.9 19.69 7.3 | 528.6 38.66 7.3 | 723.3 45.50 6.3 | 92.25 | 33 |
| Vitamin D | nmol/L | 18.1 1.81 10.0 | 37.0 2.64 7.1 | 74.2 4.44 6.0 | 6 | 33 |

Supplemental Table 1. Assay imprecision and LLOQ on the Abbott Architect.

LLOQ = lower limit of quantification.

SD = Standard deviation.

CV = Coefficient of variation.