Supplemental Table 1. Analytical assay information for biochemical parameters on the Abbott Alinity ci system.

|  |  | Alinity |  |  | ARCHITECT |
| --- | --- | --- | --- | --- | --- |
| Assay | Unit | Reference Material/Method | Method | Analytical Measuring Range | Analytical Measuring Range* |
| Alpha-1-Antitrypsin | g/L | CRM 470 | Immunoturbidimetric | 0.25 to 2.00 | 0.25 to 3.00 |
| Alpha-1-Glycoprotein | g/L | CRM 470 | Immunoturbidimetric | 0.15 to 2.50 | 0.15 to 2.50 |
| Ceruloplasmin | mg/L | ERM DA470 | Immunoturbidimetric | 20 to HC | 20 to 4000 |
| Cholinesterase | U/L | ε Hexacyanoferrate (III)/ DGKC Butyrylthiocholine Liquid/37°C | DGKC Butyrylthiocholine 37°C | 164 to 25 000 | 164 to 25 000 |
| Chloride | mmol/L | NIST/ Titration with Silver | Ion-selective electrode diluted (Indirect) | 50 to 150 | N/A |
| Complement C3 | g/L | ERM-DA470/ Reference method | Immunoturbidimetric | 0.11 to HC | 0.11 to HC |
| Complement C4 | g/L | ERM-DA470/ Reference method | Immunoturbidimetric | 0.029 to HC | 0.03 to HC |
| Glucose | mmol/L | NIST SRM 965/ ID-GC/MS | Enzymatic (Hexokinase/G-6-PDH) | 0.28 to 44.40 | N/A |
| Haptoglobin | g/L | ERM-DA470/ Reference method | Immunoturbidimetric | 0.08 to HC | 0.08 to HC |
| Immunoglobulin A | g/L | ERM-DA470/ Reference method | Immunoturbidimetric | 0.05 to HC | 0.05 to HC |
| Immunoglobulin E | IU/mL | WHO International Standard | Immunoturbidimetric | AMR: 28.4 to 1000 EMI:1000 to 10 000 RI: 15.7 to 10 000 | 25 to 1000 |
| Immunoglobulin G | g/L | ERM-DA470/ Reference method | Immunoturbidimetric | 1.09 to HC | 1.09 to HC |
| Immunoglobulin M | g/L | ERM-DA470/ Reference method | Immunoturbidimetric | 0.05 to HC | 0.05 to HC |
| Lactate | mmol/L | Reagent Grade Lactic Acid/ Titration | Lactic Acid to Pyruvate | 0.17 to 13.32 | N/A |
| Potassium | mmol/L | NIST/ Flame photometry | Ion-selective electrode diluted (Indirect) | 1.0 to 10.0 | N/A |
| Sodium | mmol/L | NIST/ Flame photometry | Ion-selective electrode diluted (Indirect) | 100 to 200 | N/A |

HC: highest calibrator level, AMR: analytical measuring range (range of values that demonstrate acceptable performance for linearity, imprecision, and bias), EMI: extended measuring interval  (extends from the upper limit of quantitation to the upper limit of quantitation x sample dilution), RI: reportable interval (extends from the lower limit of detection to the upper limit of EMI).

*extracted from: [ref 3 or 6] and/or Abbott Diagnostics historic manufacturer package inserts.

N/A: No Architect comparison included in this manuscript.

Supplemental Table 2. Analytical validation summary for biochemical parameters on the Abbott Alinity ci system.

|  | Imprecision |  | Linearity |  |  | Method Comparison |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Assay | QC Mean | QC CV (%) | Slope | Intercept | R2 | n | Slope | Intercept | R |
| Alpha-1-Antitrypsin (g/L) | Level 1: 0.99 Level 2: 1.29 Level 3: 1.56 | Level 1: 0.8 Level 2: 0.8 Level 3: 0.7 | 0.99 | 0.026 | 0.998 | 23 | 0.97 | 0.01 | 0.988 |
| Alpha-1-Glycoprotein (g/L) | Level 1: 0.51 Level 2: 0.69 Level 3: 0.84 | Level 1: 0.7 Level 2: 0.5 Level 3: 0.4 | 1.01 | 0.000 | 0.999 | 29 | 0.87 | 0.13 | 0.984 |
| Ceruloplasmin (mg/L) | Level 1: 218 Level 2: 280 Level 3: 340 | Level 1: 1.3 Level 2: 0.7 Level 3: 1.2 | - | - | - | 24 | 0.98 | 15.8 | 0.976 |
| Cholinesterase* (U/L) | Level 1: 5933 Level 2: 7777 Level 3: 9090 | Level 1: 0.4 Level 2: 0.4 Level 3: 0.4 | 1.00 | 128 | 0.999 | 6 | 0.98 | -314 | 1.000 |
| Complement C3 (g/L) | Level 1: 0.90 Level 2: 1.34 Level 3: 1.67 | Level 1: 1.3 Level 2: 0.8 Level 3: 0.8 | 0.99 | 0.011 | 0.998 | 29 | 1.07 | -0.01 | 0.995 |
| Complement C4 (g/L) | Level 1: 0.17 Level 2: 0.24 Level 3: 0.30 | Level 1: 2.9 Level 2: 2.2 Level 3: 1.9 | 1.00 | 0.001 | 1.000 | 27 | 1.05 | 0.00 | 0.991 |
| Glucose (mmol/L) | Level 1: 2.80 Level 2: 6.90 Level 3: 14.6 | Level 1: 1.1 Level 2: 0.8 Level 3: 0.9 | 1.00 | 0.458 | 0.999 | 40 | 0.98 | -0.93 | 0.995 |
| Haptoglobin (g/L) | Level 1: 0.77 Level 2: 1.03 Level 3: 1.23 | Level 1: 1.0 Level 2: 0.8 Level 3: 0.8 | 0.99 | 0.011 | 0.998 | 26 | 1.02 | 0.00 | 0.999 |
| Immunoglobulin A (g/L) | Level 1: 1.56 Level 2: 2.20 Level 3: 2.74 | Level 1: 1.4 Level 2: 1.0 Level 3: 1.0 | 0.99 | 0.032 | 0.999 | 24 | 1.04 | -0.03 | 0.999 |
| Immunoglobulin E (IU/mL) | Level 1: 70.6 Level 2: 445.8 | Level 1: 3.3 Level 2: 0.9 | 1.01 | 5.35 | 0.998 | 36 | 0.80 | 5.32 | 0.982 |
| Immunoglobulin G (g/L) | Level 1: 7.16 Level 2: 10.3 Level 3: 12.9 | Level 1: 1.2 Level 2: 0.9 Level 3: 1.0 | 1.00 | 0.201 | 0.998 | 24 | 1.07 | -0.03 | 0.999 |
| Immunoglobulin M (g/L) | Level 1: 0.61 Level 2: 0.88 Level 3:1.14 | Level 1: 1.6 Level 2: 1.0 Level 3: 1.4 | 1.00 | 0.004 | 1.000 | 21 | 1.00 | 0.00 | 0.999 |
| Lactate (mmol/L) | Level 1: 1.01 Level 2: 1.89 Level 3: 6.67 | Level 1: 1.7 Level 2: 0.5 Level 3:0.4 | 1.00 | 0.159 | 0.998 | 36 | 0.99 | 0.09 | 1.000 |
| Chloride (mmol/L) | Level 1: 79 Level 2: 95 Level 3: 108 | Level 1: 0.6 Level 2: 0.6 Level 3: 0.4 | - | - | - | 39 | 1.11 | -12.5 | 0.935 |
| Potassium (mmol/L) | Level 1: 2.8 Level 2: 4.0 Level 3: 6.3 | Level 1: 1.4 Level 2: 0.8 Level 3: 0.8 | - | - | - | 39 | 1.05 | 0.14 | 0.927 |
| Sodium (mmol/L) | Level 1: 117 Level 2: 146 Level 3: 160 | Level 1: 0.8 Level 2: 0.5 Level 3: 0.5 | - | - | - | 39 | 1.00 | -3 | 0.832 |

- not completed; method comparison statistics used ARCHITECT values as reference method; * low sample size due to lack of comparator samples.

Supplemental Table 3. Example calculation of percent verification within reference limits and confidence limits.

|  | Lower Limit | Lower Limit 90% Confidence Intervals | Upper Limit | Upper Limit 90% Confidence Intervals |
| --- | --- | --- | --- | --- |
| Variables | X | (A, B) | Y | (C,D) |
| Calculation of Percent Verification Within CALIPER ARCHITECT Reference Limits | = [number of samples falling within X and Y]/total number of samples |  |  |  |
| Calculation of Percent Verification Within CALIPER ARCHITECT Confidence Limits | = [number of samples falling within A and D]/total number of samples |  |  |  |