SUPPLEMENTAL TABLE AND FIGURE

Preservation of urine collection for metabolic evaluation of recurrent urinary stone formers.

Tomáš Šálek1,2, Pavel Musil1, Pieter Vermeersh3, Rachel Marrington4, Zeliha Gunnur Dikmen5, Radka Poláchová1, Ulrike Kipman6, Timo T. Kouri7§, and Janne Cadamuro8§,

Supplemental Table 1.

Measurement procedures on AbbottTM ArchitectTM ci16200 analyser and Lumex CAPEL-205 capillary electrophoresis. a

| Measurand |  | ## Assay application and reagent |  | ## Assay principle |  | ## Calibrator |  | ## Standard uncertainty of calibrator (k=1) |  | ## Day-to-day imprecision b, ## target value (mmol/L), CVA (%) |  | ## Relative bias (%) |  | ## Intermediate ## reproducibility, ## mean CVA of the two levels (%) |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Sodium |  | Na-CU ARCHITECT cSystems ICT, ICT Sample Diluent |  | Indirect ISE |  | ICT Urine Calibrator |  | 0.37% at 180 mmol/L |  | 63.3 mmol/L, 1.17% 207 mmol/L, 0.84% |  | 1.96% 0.24% |  | 1.01% |  |
| Potassium |  | K-CU ARCHITECT cSystems ICT, ICT Sample Diluent |  | Indirect ISE |  | ICT Urine Calibrator |  | 0.74% at 90 mmol/L |  | 29.9 mmol/L, 0.89% 133 mmol/L, 0.83% |  | 1.66% 0.14% |  | 0.86% |  |
| Chloride |  | Cl-CU ARCHITECT cSystems ICT, ICT Sample Diluent |  | Indirect ISE |  | ICT Urine Calibrator |  | 0.37% at 180 mmol/L |  | 82.0 mmol/L, 0.91% 272 mmol/L, 1.28% |  | 3.37% 1.20% |  | 1.10% |  |
| Creatinine |  | CrEnzU Creatinine (Enzymatic) |  | Spectrophotometry Creatinase |  | Clin Chem Cal |  | 0.55% at 0.48 mmol/L |  | 7.02 mmol/L, 1.48% 16.0 mmol/L, 1.50% |  | 1.24% 1.23% |  | 1.49% |  |
| Calcium |  | CaCU Calcium |  | Spectrophotometry Arsenazo III |  | Multiconstituent Calibrator |  | 0.2% at 3 mmol/L |  | 1.53 mmol/L, 1.94% 4.21 mmol/L, 1.84% |  | 4.08% 6.05% |  | 1.89% |  |
| Magnesium |  | MAGU Magnesium |  | Spectrophotometry Isocitrate dehydrogenase |  | Multiconstituent Calibrator |  | 0.21% at 1.93 mmol/L |  | 3.09 mmol/L, 2.28% 13.1 mmol/L, 2.14% |  | 5.96% 2.38% |  | 2.21% |  |
| Phosphate |  | Phos-U Phosphorus |  | Spectrophotometry Phosphomolybdate |  | Multiconstituent Calibrator |  | 0.17% at 2.54 mmol/L |  | 9.00 mmol/L, 2.16% 26.8 mmol/L, 1.18% |  | 1.22% 0.82% |  | 1.67% |  |
| Uric acid |  | Uric2-U Uric Acid2 |  | Spectrophotometry Uricase |  | Consolidated Chemistry Calibrator |  | 0.22% at 0.55 mmol/L |  | 0.702 mmol/L, 2.16% 1.21 mmol/L, 1.58% |  | 4.56% 2.84% |  | 1.87% |  |
| Oxalate |  | Capillary electrophoresis |  | Indirect UV detection |  | Ammonium oxalate monohydrate |  | 1.2% at 1.00 mmol/L |  | 0.201 mmol/L (3.87)b 0.484 mmol/L (3.81)b |  | 0.97% 2.95% |  | 3.84% |  |
| Citrate |  | Capillary electrophoresis |  | Indirect UV detection |  | Sodium citrate dihydrate |  | 2.0% at 6.00 mmol/L |  | 1.200 mmol/L (4.02)b 3.00 mmol/L (3.47)b |  | 2.06% 2.69% |  | 3.75% |  |

a Abbreviations in the names of materials of the AbbottTM ArchitectTM analyzers represent the commercial names of those products.
b Randox Assayed Urine Controls (Crumlin, UK; on AbbottTM ArchitectTM), or in-house solutions (from Penta s.r.o, Prague, Czech Republic; on Lumex CAPEL-205 capillary electrophoresis) were measured once on each working day during the study, to calculate an intermediate reproducibility of 580 consecutive measurements. CVA = coefficient of analytical variation.

Supplemental Figure 1: Typical oxalate crystals in the urine sample of one patient.  Both octahedral calcium oxalate-dihydrate (COD) and round-shaped calcium oxalate monohydrate (COM) crystals were detected quantitatively by the iQ200 instrument, as confirmed by visual microscopy.