SUPPLEMENTAL TABLE 1

Supplemental Table S1. Overview of studies evaluating the performance of line/dot blots for the identification of PNS antibodies.

| Reference and type of study | Number of samples Type of patients | Technique | (Comparative) results | Clinical associations | Conclusion authors | Comments/ study limitations |
| --- | --- | --- | --- | --- | --- | --- |
| Ali et al. Frontiers in Neurology 2025 [15] Single-center retrospective study Comparison between different test methods (LB and IIF cerebellum) in relation to clinical information | Patient samples: N=2081 Consecutive patient samples analyzed by lab for PNS Ab Detailed analysis/clinical review of N= 91 LB+ positive samples (from 69 patients) - 35 serum only - 15 CSF only - 16 combination - 3 serial serum Healthy donor samples: N= 71 sera | EUROLINE PNS 12 Ag (Euroimmun) (serum dilution 1:10; CSF dilution 1:4) + automated reading Rat cerebellum IIF Inova Diagnostics (dil. 1:1 and 1:50, optional 1:100) Tested reactivities: CV2, GAD65; Hu, Ma2/Ta, recoverin, Ri, Tr, Sox-1, Titine, Yo, Zic4, amphiphysine | 4.4% samples (91/2081, 69 patients) positive for PNS Ab* 6/69 patients (9%) multiple reactive Concordance on LB between serum and CSF Majority of LB+ samples had no (typical) IIF pattern (90%) | High FP rate on LB: - PPV%= 9/69 - Spec%=93% Higher clinical relevance in case of: - Higher titers (Hu, Tr Amphiphysine) - Confirmed by IIF | LB has more FP than TP results Use should be limited to those with high clinical suspicion IIF only is insufficient for screening | *Anti-recoverin and -titine positive samples excluded from the analysis (n= 48) *weak bands (6-10U) were not considered positive *samples originating from the same patients were not excluded |
| Ruiz-Garcia et al. Frontiers in Immunology 2020 [16] Single-center retrospective study Comparison between different test methods (LB and IIF cerebellum) in relation to clinical information | Patient serum samples N=2437 Consecutive patient samples analyzed by lab for PNS Ab Detailed analysis/clinical review of N= 493 samples analyzed with both IIF and LB | EUROLINE PNS 12 Ag (Euroimmun) (serum dilution 1:100; CSF dilution 1:4) + automated reading In-house developed rat cerebellum (dil. 1:500) Tested reactivities: CV2, GAD65; Hu, Ma2/Ta, recoverin, Ri, Tr, Sox-1, Titine, Yo, Zic4, amphiphysin | 22% (96/439) positivity for PNS Ab* 18/96 patients (19%) multiple reactive Concordance LB and IIF: - 52% (n=50) discordant samples (LB+ and IIF-). - High concordance (93%) in case high positive result on LB | Concordant results more likely to have: - PNS: 91% (LB+/IIF+) vs. 8% (LB+/IIF-) - Cancer: 91% (LB+/IIF+) vs. 30% (LB+/IIF-) High positive LB results are highly predictive for PNS (PPV 86%) PPV% of LB+>10U for PNS (irrespective of IIF- concordant with IIF): Hu: 93%-100%, Sox1:53%-85%; CV2: 50%-80%; Yo1: 29%-100%; Zic: 0%-0%; other: 59%-83% | Patients with suspected PNS should have dual testing Anti-ZIC4 and -Yo only detected on LB are likely FP For other Ab, strong intensity on LB correlates with IIF+ and PNS diagnosis LB results <50U, not confirmed on IIF are les predictive for PNS | *Selection bias: no systematic analysis with both techniques - LB+ - On physician request - Confirmation of results from other lab *Anti-recoverin, -titine and GAD65 positive samples excluded from the analysis * isolated neurological symptoms were not considered PNS regardless of the presence of cancer or PNS Ab |
| Déchelotte et al., Neurol Neuroimmunol Neuroinflam 2020 [14] Single-center retrospective study Comparison between different test methods (2 LBs, IIF cerebellum, Western blotting and ell-based assays) in relation to clinical information | Patient serum samples: N=5300 Consecutive patient samples analyzed by lab for PNS Ab Detailed analysis/clinical review of N= 314 LB+ positive samples | PNS+2 blot (LB1) (Ravo Diagnostica) + automated reading* EUROLINE PNS 12 Ag (LB2) (Euroimmun) + automated reading* Confirmatory testing by: - In-house developed rat cerebellum (dil. 1:100) - WB for anti-CV2 and amphiphysine - CBs for all other PNS Ab Tested reactivities: CV2, GAD65; Hu, Ma2/Ta, recoverin, Ri, Tr, Sox-1, Titine, Yo, Zic4, amphiphysine, Ma1 | 6.2% (330/5300) positivity for PNS Ab on one of both LBs (LB1: 7.4%; LB2: 5.1%) 15.3 % (LB1) and 21.5% (LB2%) multiple reactive Confirmation rate (overall 33%): - LB1: 36.7% of sera, with varying confirmation rate per PNS Ab (0-88%) - LB2: 30.1% of sera, with varying confirmation rate per PNS Ab (0-65%); weak positive LB results (8-14U) are never confirmed; relation between signal intensity and final confirmation was variable amongst Ab | Confirmed LB+ results are likely to have: - PNS: 100% - Cancer: 86.2% All non-confirmed LB+ results had no PNS | Less than 1/3 of LB+ results can be confirmed using alternative techniques Positive LB+ results should be correlated with clinical features, and confirmation are reference centers are needed | *serum dilution not specified but referred to manufacturer’s instructions *Anti-GAD65, -titine and -recoverine excluded from the analysis |

Abbreviations: LB= lineblot; WB: Western Blot; antibodies: Ab; IIF; indirect immunofluorescence analysis on cerebellum;  PNS: paraneoplastic syndrome; FP: false positive; TP: true positive; PPV: positive predictive value; CSF: cerebrospinal fluid